## Slide 1: Facilitating testing and monitoring of number entry systems in medical devices
Abigail Cauchi, Christian Colombo,
Mark Micallef & Gordon Pace

## Slide 2
Notes: 5 Key number entry interfaces are used in critical situations such as hospital intensive care units.

## Slide 4
Same input key sequence:
  Left
  Up 5 times
  Left
  Down
Zimed AD
BBraun Infusomat Space
Notes: Entering the same volume in both devices with the same interface gives different results.

## Slide 5
Under the UK Health & Safety At Work Act (1974) and under similar legislation in other countries, devices should be designed to reduce risk to be
As Low As Reasonably Practical, ALARP.
Notes: Under the UK Health & Safety At Work Act (1974) and under similar legislation in other countries, devices should be designed to reduce risk to be As Low As Reasonably Practical, ALARP..

## Slide 6: UI principles
No button presses with no effects
Understandable error messages
Reduce chances of (big) mistakes
Etc

## Slide 7: Cursor wraparound
Notes: An example scenario that we will want to verify (talking about cursor wraparound)

## Slide 8: The specification
Coming from UI experts, psychologists, etc
Implemented by software developers

## Slide 9: The specification
Coming from UI experts, psychologists, etc
Implemented by software developers
The problem is not new!
Customers already specify systems

## Slide 10: The specification
Coming from UI experts, psychologists, etc
Implemented by software developers
The problem is not new!
Customers already specify systems
Gherkin

## Slide 11: Gherkin Feature
Notes: Gherkin feature written in the .feature file what would have been drawn up during the three amigos meeting

## Slide 12: Automated Testing
How can we use Gherkin to execute tests?

## Slide 13: Gherkin output
Notes: When the feature file is run, this is what is output in the console.
This can be copied and pasted in the StepDefinitions.java file and filled in by the tester.

## Slide 14: StepDefinitions.java
Notes: This is the implemented StepDefinitions.java file.

## Slide 15: Automated Testing
How can we use Gherkin to execute tests?
Execute the scenarios, invoking the steps

## Slide 16: Logs from Running the Gherkin Feature
Notes: When the .feature file is run, this is the output in the console.

## Slide 17: Runtime Monitoring
Notes: + Motivation
We want to make sure that things are running well at runtime and runtime monitoring lets us do that.
This isn’t widely adopted by industry, and one of the challenges is finding the expertise

## Slide 18: Runtime Monitoring
Can we use the tests to automatically generate monitors?

## Slide 19: Gherkin Feature
Notes: Gherkin feature written in the .feature file what would have been drawn up during the three amigos meeting

## Slide 20: Gherkin Feature
Precondition
Notes: Gherkin feature written in the .feature file what would have been drawn up during the three amigos meeting

## Slide 21: Gherkin Feature
Precondition
Event
Notes: Gherkin feature written in the .feature file what would have been drawn up during the three amigos meeting

## Slide 22: Gherkin Feature
Precondition
Event
Assertion
Notes: Gherkin feature written in the .feature file what would have been drawn up during the three amigos meeting

## Slide 23: Aspect Oriented Programming Implementation
Event
Notes: We implemented

## Slide 24: Aspect Oriented Programming Implementation
Event
Precondition
Notes: We implemented

## Slide 25: Aspect Oriented Programming Implementation
Event
Assertions
Precondition
Notes: We implemented

## Slide 26: Aspect Oriented Programming Implementation
Event
Precondition
+ Setup
Assertions
Crucial to match the precondition
with the corresponding assertion
Crucial to match the precondition
with the corresponding assertion
Crucial to match the precondition
with the corresponding assertion
Notes: We implemented

## Slide 27: Things to note
Event
Assertions
Easy to extract
Precondition
Easy to extract
Notes: We implemented

## Slide 28: Things to note
Event
Precondition
+ Setup
Assertions
Not straight forward:
Going from action to condition!
Notes: We implemented

## Slide 29: One approach
Notes: We implemented

## Slide 30: One approach
Notes: We implemented

## Slide 31: One approach
How does this interact with other methods?
Eg: Are there other methods which change the cursor position?
Does it affect other parts of the system state?
Notes: Two options: either be very strict and turn the boolean off in case of other method call, or assume that through encapsulation, cursor position can only be changed through this method call

## Slide 32: One approach
How does this interact with other methods?
Eg: Are there other methods which change the cursor position?
Does it affect other parts of the system state?
Static analysis to identify such methods and reset the flag if detected
Notes: Two options: either be very strict and turn the boolean off in case of other method call, or assume that through encapsulation, cursor position can only be changed through this method call

## Slide 33: Left() method implementation
Notes: Left() implementation in the system

## Slide 34: Aspect Oriented Programming Implementation

## Slide 35: Discussion
Resulting monitors can be too specific
Can generalise using tester input but we want to automate

## Slide 36: Options to consider
Observing tests and extract invariants
  Automatically deduce pre-post conditions:
    cursor == 0
    Left
    cursor == 0

## Slide 37: Options to consider
Observing tests and extract invariants
  Automatically deduce pre-post conditions:
    cursor == 0
    Left
    cursor == 0
How much can/should you generalise from observing a few tests?

## Slide 38: Filtering
Filtering generated invariants
vs
Filtering resulting monitor alerts
Attempting to model check the
invariant to check if it is true
Attempting to choose variables to consider for invariant extraction

## Slide 39: Filtering
Filtering generated invariants
vs
Filtering resulting monitor alerts
Showing the monitor results in priority order
Asking user for feedback

## Slide 40: Conclusions
Numerous challenges when going from tests to monitors
Right mix of techniques to obtain all available information: Static analysis, dynamic analysis on test + system description, test + system execution.
Right mix of user input and automation